## OD600 reference point
| LUDOX-HS40 | H2O
Replicate 1 | 0.046 | 0.035 | Enter Abs600 absorbance measurements into blue cells
Replicate 2 | 0.045 | 0.036 | Gold cells are calculated
Replicate 3 | 0.046 | 0.036
Replicate 4 | 0.047 | 0.037
Arith. Mean | 0.046 | 0.036
Corrected Abs600 | 0.01 |  | Corrected value is particle-only contribution
Reference OD600 | 0.0425 |  | Reference value is for 100uL of LUDOX-HS40 in a well of a standard 96-well flat-bottom plate
OD600/Abs600 | 4.25 |  | Corrected value = scaling factor * measured value

## Fluorescein standard curve
uM Fluorescein | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625 | 0.78125 | 0.39062 | 0.19531 | 0.097656 | 0.048828 | 0
Replicate 1 | 1753 | 1091 | 612.6 | 345 | 157.3 | 77.01 | 31.56 | 13.93 | 6.117 | 2.58 | 1.055 | 0.0408 | Enter fluorescence measurements into blue cells
Replicate 2 | 1861 | 949 | 411.6 | 261 | 107.9 | 42.05 | 19.91 | 9.346 | 3.243 | 1.872 | 0.6322 | 0.0342 | Gold cells are calculated
Replicate 3 | 1786 | 1283 | 342.5 | 125.3 | 65.8 | 26.04 | 11.75 | 5.961 | 2.926 | 1.249 | 0.6074 | 0.0399
Replicate 4 | 1851 | 957.2 | 575.3 | 209.4 | 97.26 | 45.93 | 11.73 | 6.949 | 3.559 | 0.7711 | 0.4975 | 0.038 | Values measured are fluorescence from 100uL of X uM fluorescein solution
Arith. Mean | 1812.75 | 1070.05 | 485.5 | 235.175 | 107.065 | 47.7575 | 18.7375 | 9.0465 | 3.9612 | 1.618 | 0.69803 | 0.038225
Arith. Std.Dev. | 51.8869 | 156.1781 | 129.2654 | 92.1327 | 37.9613 | 21.3172 | 9.3759 | 3.5524 | 1.4602 | 0.78387 | 0.24508 | 0.0029262
Values should form a straight line on both linear and log scale
Slope should be 1:1
Common problems:
* Consistent pipetting error --> log graph is a straight line but not 1:1 slope
* Oversaturated detector --> low concentrations linear, but high concentrations saturate or fall
uM Fluorescein/a.u. | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625 | 0.78125 | 0.39062 | 0.19531 | 0.097656 | 0.048828
Mean um Fluorescein/a.u. | 0.027582 | 0.023363 | 0.025747 | 0.026576 | 0.029188 | 0.032717 | 0.041694 | 0.04318 | 0.049306 | 0.060355 | 0.069952
Mean of med-high levels: |  | 0.027518
Final scaling level determined from medium-high points likely to be less impacted by saturation or pipetting error
If needed, you can shift which points are used, but it is likely better to correct instrument settings and protocol.

## Raw Plate Reader Measurements
Raw Plate Readings |  | If you followed the recommended plate layout:
Copy fluorescence and Abs600 measurements from your plate reader into blue cells
They will automatically propagate into the correct locations in the Fluorescence Measurement Sheet
Fluorescence Raw Readings: | [9 empty] | Abs600 Raw Readings:
Hour 0: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 0: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 3.985 | 5.906 | 8.287 | 9.338 | 3.82 | 11.27 | 4.172 | 3.708 | 3.636 | Colony 1, Replicate 1 | 0.1 | 0.066 | 0.301 | 0.075 | 0.17 | 0.208 | 0.304 | 0.152 | 0.05
Colony 1, Replicate 2 | 3.771 | 5.552 | 7.289 | 9.078 | 3.939 | 10.87 | 4.037 | 3.639 | 3.608 | Colony 1, Replicate 2 | 0.096 | 0.065 | 0.093 | 0.13 | 0.072 | 0.161 | 0.203 | 0.144 | 0.05
Colony 1, Replicate 3 | 3.815 | 5.242 | 5.828 | 8.199 | 3.67 | 11.18 | 3.089 | 3.617 | 3.629 | Colony 1, Replicate 3 | 0.135 | 0.064 | 0.17 | 0.073 | 0.17 | 0.136 | 0.06 | 0.131 | 0.049
Colony 1, Replicate 4 | 3.63 | 5.019 | 6.256 | 6.758 | 3.672 | 12.69 | 4.082 | 3.724 | 3.599 | Colony 1, Replicate 4 | 0.095 | 0.108 | 0.15 | 0.059 | 0.151 | 0.122 | 0.151 | 0.074 | 0.05
Colony 2, Replicate 1 | 3.879 | 5.605 | 7.153 | 6.924 | 3.794 | 8.107 | 3.952 | 3.645 | 3.808 | Colony 2, Replicate 1 | 0.075 | 0.067 | 0.076 | 0.103 | 0.075 | 0.11 | 0.145 | 0.102 | 0.048
Colony 2, Replicate 2 | 3.853 | 5.474 | 7.155 | 6.45 | 3.732 | 8.077 | 3.936 | 3.671 | 3.819 | Colony 2, Replicate 2 | 0.071 | 0.061 | 0.069 | 0.077 | 0.068 | 0.088 | 0.114 | 0.082 | 0.053
Colony 2, Replicate 3 | 3.841 | 4.819 | 6.884 | 6.331 | 3.69 | 8.141 | 4.059 | 3.886 | 3.783 | Colony 2, Replicate 3 | 0.066 | 0.064 | 0.07 | 0.058 | 0.078 | 0.113 | 0.065 | 0.077 | 0.048
Colony 2, Replicate 4 | 3.714 | 4.793 | 7.1 | 6.036 | 3.765 | 8.791 | 4.043 | 3.755 | 3.845 | Colony 2, Replicate 4 | 0.066 | 0.061 | 0.107 | 0.071 | 0.069 | 0.071 | 0.11 | 0.07 | 0.05
Hour 2: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 2: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 3.935 | 9.619 | 7.927 | 17.02 | 4.145 | 12.43 | 6.417 | 3.616 | 3.767 | Colony 1, Replicate 1 | 0.419 | 0.221 | 0.249 | 0.176 | 0.382 | 0.204 | 0.36 | 0.206 | 0.052
Colony 1, Replicate 2 | 3.977 | 9.845 | 7.427 | 14.29 | 3.822 | 11.73 | 5.995 | 4.102 | 3.869 | Colony 1, Replicate 2 | 0.207 | 0.158 | 0.179 | 0.175 | 0.249 | 0.149 | 0.274 | 0.225 | 0.048
Colony 1, Replicate 3 | 3.669 | 8.415 | 7.174 | 13.53 | 3.865 | 11.67 | 6.04 | 3.682 | 3.756 | Colony 1, Replicate 3 | 0.242 | 0.162 | 0.162 | 0.159 | 0.257 | 0.181 | 0.201 | 0.181 | 0.05
Colony 1, Replicate 4 | 3.683 | 8.481 | 6.333 | 15.27 | 3.808 | 13.63 | 6.064 | 3.955 | 3.677 | Colony 1, Replicate 4 | 0.195 | 0.126 | 0.177 | 0.183 | 0.229 | 0.127 | 0.206 | 0.181 | 0.051
Colony 2, Replicate 1 | 3.774 | 10.01 | 7.626 | 12.22 | 3.978 | 10.73 | 6.392 | 3.803 | 3.639 | Colony 2, Replicate 1 | 0.212 | 0.155 | 0.123 | 0.146 | 0.209 | 0.139 | 0.181 | 0.158 | 0.048
Colony 2, Replicate 2 | 3.582 | 9.419 | 7.287 | 11.69 | 3.851 | 10.44 | 6.139 | 3.901 | 3.768 | Colony 2, Replicate 2 | 0.2 | 0.131 | 0.128 | 0.111 | 0.175 | 0.086 | 0.185 | 0.132 | 0.049
Colony 2, Replicate 3 | 3.273 | 7.976 | 7.072 | 11.9 | 3.866 | 10.69 | 6.22 | 3.846 | 3.78 | Colony 2, Replicate 3 | 0.171 | 0.129 | 0.123 | 0.157 | 0.203 | 0.116 | 0.191 | 0.162 | 0.049
Colony 2, Replicate 4 | 3.634 | 8.518 | 7.061 | 11.97 | 3.536 | 12.67 | 6.096 | 3.914 | 3.652 | Colony 2, Replicate 4 | 0.186 | 0.094 | 0.109 | 0.119 | 0.198 | 0.086 | 0.184 | 0.125 | 0.049
Hour 4: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 4: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 4.106 | 15.96 | 8.389 | 23.84 | 4.397 | 12.25 | 7.217 | 4.223 | 3.902 | Colony 1, Replicate 1 | 0.386 | 0.202 | 0.183 | 0.238 | 0.313 | 0.233 | 0.343 | 0.266 | 0.053
Colony 1, Replicate 2 | 3.966 | 14.43 | 8.387 | 23.58 | 4.284 | 11.89 | 7.029 | 4.154 | 4.057 | Colony 1, Replicate 2 | 0.334 | 0.164 | 0.138 | 0.193 | 0.257 | 0.162 | 0.252 | 0.253 | 0.051
Colony 1, Replicate 3 | 3.965 | 15.06 | 8.374 | 24.47 | 4.224 | 1.99 | 6.908 | 3.93 | 3.841 | Colony 1, Replicate 3 | 0.32 | 0.169 | 0.127 | 0.221 | 0.27 | 0.077 | 0.118 | 0.232 | 0.052
Colony 1, Replicate 4 | 3.796 | 16.85 | 8.47 | 24.18 | 4.114 | 11.74 | 6.907 | 3.988 | 3.828 | Colony 1, Replicate 4 | 0.323 | 0.176 | 0.135 | 0.22 | 0.332 | 0.083 | 0.213 | 0.236 | 0.055
Colony 2, Replicate 1 | 3.712 | 15.96 | 7.627 | 19.62 | 4.426 | 11.69 | 7.026 | 4.245 | 3.599 | Colony 2, Replicate 1 | 0.265 | 0.232 | 0.114 | 0.185 | 0.218 | 0.089 | 0.205 | 0.181 | 0.049
Colony 2, Replicate 2 | 3.741 | 16.19 | 7.939 | 18.82 | 3.922 | 11.07 | 6.85 | 4.192 | 3.801 | Colony 2, Replicate 2 | 0.271 | 0.176 | 0.099 | 0.182 | 0.228 | 0.079 | 0.189 | 0.178 | 0.051
Colony 2, Replicate 3 | 3.58 | 16.49 | 7.63 | 19.69 | 4.078 | 12.45 | 6.502 | 4.024 | 3.974 | Colony 2, Replicate 3 | 0.307 | 0.206 | 0.13 | 0.192 | 0.247 | 0.089 | 0.194 | 0.19 | 0.05
Colony 2, Replicate 4 | 3.58 | 16.96 | 7.988 | 20.21 | 3.93 | 12.98 | 6.833 | 4.266 | 3.675 | Colony 2, Replicate 4 | 0.289 | 0.163 | 0.116 | 0.2 | 0.276 | 0.081 | 0.196 | 0.184 | 0.056
Hour 6: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 6: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 4.374 | 29.19 | 9.903 | 35.49 | 4.552 | 13.73 | 7.989 | 4.282 | 3.849 | Colony 1, Replicate 1 | 0.413 | 0.315 | 0.341 | 0.294 | 0.502 | 0.24 | 0.455 | 0.365 | 0.05
Colony 1, Replicate 2 | 4.253 | 24.21 | 9.307 | 34.2 | 4.491 | 13.48 | 8.519 | 4.116 | 3.714 | Colony 1, Replicate 2 | 0.383 | 0.286 | 0.238 | 0.361 | 0.425 | 0.191 | 0.267 | 0.364 | 0.048
Colony 1, Replicate 3 | 4.099 | 28.59 | 9.974 | 34.61 | 4.471 | 13.24 | 8.006 | 4.094 | 3.603 | Colony 1, Replicate 3 | 0.394 | 0.302 | 0.21 | 0.289 | 0.426 | 0.171 | 0.278 | 0.36 | 0.056
Colony 1, Replicate 4 | 4.364 | 27.84 | 9.856 | 33.49 | 4.705 | 13.3 | 7.904 | 4.008 | 3.7 | Colony 1, Replicate 4 | 0.395 | 0.312 | 0.208 | 0.331 | 0.313 | 0.151 | 0.342 | 0.339 | 0.051
Colony 2, Replicate 1 | 4.123 | 26.78 | 9.45 | 27.4 | 4.297 | 16.8 | 7.333 | 4.113 | 3.677 | Colony 2, Replicate 1 | 0.362 | 0.342 | 0.176 | 0.347 | 0.363 | 0.173 | 0.328 | 0.317 | 0.047
Colony 2, Replicate 2 | 4.096 | 27.11 | 9.541 | 27.24 | 4.397 | 16.81 | 7.335 | 4.12 | 4.067 | Colony 2, Replicate 2 | 0.341 | 0.323 | 1.159 | 0.3 | 0.258 | 0.133 | 0.288 | 0.298 | 0.055
Colony 2, Replicate 3 | 4.029 | 26.46 | 9.245 | 26.3 | 4.202 | 16.84 | 7.997 | 4.043 | 3.682 | Colony 2, Replicate 3 | 0.334 | 0.346 | 0.156 | 0.268 | 0.334 | 0.163 | 0.276 | 0.318 | 0.048
Colony 2, Replicate 4 | 4.33 | 27.21 | 9.339 | 29.43 | 4.268 | 16.73 | 7.476 | 4.578 | 3.702 | Colony 2, Replicate 4 | 0.349 | 0.277 | 0.143 | 0.294 | 0.323 | 0.12 | 0.299 | 0.289 | 0.05
Plate pattern:
| A1 | A2 | A3 | A4 | A5 | A6 | A7 | A8 | A9
| B1 | B2 | B3 | B4 | B5 | B6 | B7 | B8 | B9
| C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9
| D1 | D2 | D3 | D4 | D5 | D6 | D7 | D8 | D9
| E1 | E2 | E3 | E4 | E5 | E6 | E7 | E8 | E9
| F1 | F2 | F3 | F4 | F5 | F6 | F7 | F8 | F9
| G1 | G2 | G3 | G4 | G5 | G6 | G7 | G8 | G9
| H1 | H2 | H3 | H4 | H5 | H6 | H7 | H8 | H9

## Fluorescence Measurement
Unit Scaling Factors: | These are imported from the prior two sheets | [5 empty] | Enter fluorescence and Abs600 measurements into blue cells on "Raw Plate Reader Measurements"
OD600/Abs600 | 4.25 | [5 empty] | They will be copied into the green cells on this sheet.
uM Fluorescein/a.u. | 0.027518 | [5 empty] | Gold cells are calculated
If you have more replicates, unhide the extra columns
Experimental Values: | Raw Abs600 | [5 empty] | Raw Fluorescence
Sample set: | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6
Blank media | 0.05 | 0.05 | 0.049 | 0.05 |  |  | 3.636 | 3.608 | 3.629 | 3.599
Blank mean: | 0.04975 | [5 empty] | 3.618 | [5 empty] | OD - Background | [5 empty] | Fluorescence - Background | [5 empty] | uM Fluorescein / OD600 | [5 empty] | Summary Statistics |  |  |  | Ln uM FITC / OD600
Hour 0: | [12 empty] | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Arith. Mean | Arith. Std.Dev. | Geo. Mean | Geo. Std. Dev. | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6
Negative Control (Colony 1) | 0.1 | 0.096 | 0.135 | 0.095 |  |  | 3.985 | 3.771 | 3.815 | 3.63 |  |  | 0.05025 | 0.04625 | 0.08525 | 0.04525 | --- | --- | 0.367 | 0.153 | 0.197 | 0.012 | --- | --- | 0.047289 | 0.02142 | 0.014962 | 0.0017171 | --- | --- | 0.021347 | 0.019141 | 0.012701 | 4.1289 | -3.0515 | -3.8434 | -4.2022 | -6.3671 | --- | ---
Negative Control (Colony 2) | 0.075 | 0.071 | 0.066 | 0.066 |  |  | 3.879 | 3.853 | 3.841 | 3.714 |  |  | 0.02525 | 0.02125 | 0.01625 | 0.01625 | --- | --- | 0.261 | 0.235 | 0.223 | 0.096 | --- | --- | 0.066928 | 0.071605 | 0.088855 | 0.038252 | --- | --- | 0.06641 | 0.021007 | 0.063529 | 1.4321 | -2.7041 | -2.6366 | -2.4207 | -3.2636 | --- | ---
Positive Control (Colony 1) | 0.066 | 0.065 | 0.064 | 0.108 |  |  | 5.906 | 5.552 | 5.242 | 5.019 |  |  | 0.01625 | 0.01525 | 0.01425 | 0.05825 | --- | --- | 2.288 | 1.934 | 1.624 | 1.401 | --- | --- | 0.91166 | 0.82114 | 0.73791 | 0.15573 | --- | --- | 0.65661 | 0.34138 | 0.54157 | 2.3057 | -0.092484 | -0.19706 | -0.30393 | -1.8596 | --- | ---
Positive Control (Colony 2) | 0.067 | 0.061 | 0.064 | 0.061 |  |  | 5.605 | 5.474 | 4.819 | 4.793 |  |  | 0.01725 | 0.01125 | 0.01425 | 0.01125 | --- | --- | 1.987 | 1.856 | 1.201 | 1.175 | --- | --- | 0.74583 | 1.0682 | 0.54571 | 0.67627 | --- | --- | 0.759 | 0.2222 | 0.73637 | 1.3234 | -0.29326 | 0.065986 | -0.60567 | -0.39117 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.301 | 0.093 | 0.17 | 0.15 |  |  | 8.287 | 7.289 | 5.828 | 6.256 |  |  | 0.25125 | 0.04325 | 0.12025 | 0.10025 | --- | --- | 4.669 | 3.671 | 2.21 | 2.638 | --- | --- | 0.12032 | 0.54958 | 0.119 | 0.17038 | --- | --- | 0.23982 | 0.20789 | 0.19135 | 2.0603 | -2.1176 | -0.5986 | -2.1286 | -1.7697 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.076 | 0.069 | 0.07 | 0.107 |  |  | 7.153 | 7.155 | 6.884 | 7.1 |  |  | 0.02625 | 0.01925 | 0.02025 | 0.05725 | --- | --- | 3.535 | 3.537 | 3.266 | 3.482 | --- | --- | 0.87195 | 1.1897 | 1.0443 | 0.39381 | --- | --- | 0.87494 | 0.34605 | 0.80818 | 1.642 | -0.13702 | 0.1737 | 0.043342 | -0.93189 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.075 | 0.13 | 0.073 | 0.059 |  |  | 9.338 | 9.078 | 8.199 | 6.758 |  |  | 0.02525 | 0.08025 | 0.02325 | 0.00925 | --- | --- | 5.72 | 5.46 | 4.581 | 3.14 | --- | --- | 1.4668 | 0.44053 | 1.2758 | 2.198 | --- | --- | 1.3453 | 0.72232 | 1.1602 | 1.9849 | 0.38307 | -0.81977 | 0.24354 | 0.78753 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.103 | 0.077 | 0.058 | 0.071 |  |  | 6.924 | 6.45 | 6.331 | 6.036 |  |  | 0.05325 | 0.02725 | 0.00825 | 0.02125 | --- | --- | 3.306 | 2.832 | 2.713 | 2.418 | --- | --- | 0.40199 | 0.67291 | 2.1293 | 0.73677 | --- | --- | 0.98523 | 0.77637 | 0.80711 | 2.0129 | -0.91133 | -0.39614 | 0.75577 | -0.30549 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.17 | 0.072 | 0.17 | 0.151 |  |  | 3.82 | 3.939 | 3.67 | 3.672 |  |  | 0.12025 | 0.02225 | 0.12025 | 0.10125 | --- | --- | 0.202 | 0.321 | 0.052 | 0.054 | --- | --- | 0.010877 | 0.093413 | 0.0027999 | 0.0034533 | --- | --- | 0.027636 | 0.044004 | 0.0099557 | 4.9896 | -4.5211 | -2.3707 | -5.8782 | -5.6684 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.075 | 0.068 | 0.078 | 0.069 |  |  | 3.794 | 3.732 | 3.69 | 3.765 |  |  | 0.02525 | 0.01825 | 0.02825 | 0.01925 | --- | --- | 0.176 | 0.114 | 0.072 | 0.147 | --- | --- | 0.045132 | 0.040446 | 0.016502 | 0.049445 | --- | --- | 0.037881 | 0.014719 | 0.034935 | 1.6598 | -3.0982 | -3.2078 | -4.1043 | -3.0069 | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.208 | 0.161 | 0.136 | 0.122 |  |  | 11.27 | 10.87 | 11.18 | 12.69 |  |  | 0.15825 | 0.11125 | 0.08625 | 0.07225 | --- | --- | 7.652 | 7.252 | 7.562 | 9.072 | --- | --- | 0.31309 | 0.42208 | 0.56769 | 0.81301 | --- | --- | 0.52897 | 0.21619 | 0.49695 | 1.5043 | -1.1613 | -0.86257 | -0.56618 | -0.20701 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.11 | 0.088 | 0.113 | 0.071 |  |  | 8.107 | 8.077 | 8.141 | 8.791 |  |  | 0.06025 | 0.03825 | 0.06325 | 0.02125 | --- | --- | 4.489 | 4.459 | 4.523 | 5.173 | --- | --- | 0.48242 | 0.75481 | 0.46302 | 1.5762 | --- | --- | 0.81912 | 0.52202 | 0.71799 | 1.7665 | -0.72894 | -0.28129 | -0.76999 | 0.45503 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.304 | 0.203 | 0.06 | 0.151 |  |  | 4.172 | 4.037 | 3.089 | 4.082 |  |  | 0.25425 | 0.15325 | 0.01025 | 0.10125 | --- | --- | 0.554 | 0.419 | -0.529 | 0.464 | --- | --- | 0.014108 | 0.017703 | -0.33417 | 0.029673 | --- | --- | -0.068171 | 0.17746 | #NUM! | #NUM! | -4.261 | -4.034 | #NUM! | -3.5175 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.145 | 0.114 | 0.065 | 0.11 |  |  | 3.952 | 3.936 | 4.059 | 4.043 |  |  | 0.09525 | 0.06425 | 0.01525 | 0.06025 | --- | --- | 0.334 | 0.318 | 0.441 | 0.425 | --- | --- | 0.022705 | 0.032047 | 0.18724 | 0.045673 | --- | --- | 0.071916 | 0.077459 | 0.049945 | 2.5245 | -3.7852 | -3.4406 | -1.6754 | -3.0862 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.152 | 0.144 | 0.131 | 0.074 |  |  | 3.708 | 3.639 | 3.617 | 3.724 |  |  | 0.10225 | 0.09425 | 0.08125 | 0.02425 | --- | --- | 0.09 | 0.021 | -0.001 | 0.106 | --- | --- | 0.0056992 | 0.0014427 | -7.9691e-05 | 0.028303 | --- | --- | 0.0088412 | 0.013203 | #NUM! | #NUM! | -5.1674 | -6.5413 | #NUM! | -3.5648 | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.102 | 0.082 | 0.077 | 0.07 |  |  | 3.645 | 3.671 | 3.886 | 3.755 |  |  | 0.05225 | 0.03225 | 0.02725 | 0.02025 | --- | --- | 0.027 | 0.053 | 0.268 | 0.137 | --- | --- | 0.0033459 | 0.010641 | 0.06368 | 0.043805 | --- | --- | 0.030368 | 0.028341 | 0.017752 | 3.8702 | -5.7 | -4.5431 | -2.7539 | -3.128 | --- | ---
Hour 2:
Negative Control (Colony 1) | 0.419 | 0.207 | 0.242 | 0.195 |  |  | 3.935 | 3.977 | 3.669 | 3.683 |  |  | 0.36925 | 0.15725 | 0.19225 | 0.14525 | --- | --- | 0.317 | 0.359 | 0.051 | 0.065 | --- | --- | 0.0055587 | 0.014782 | 0.0017177 | 0.0028975 | --- | --- | 0.006239 | 0.0059176 | 0.0044969 | 2.5281 | -5.1924 | -4.2143 | -6.3668 | -5.8439 | --- | ---
Negative Control (Colony 2) | 0.212 | 0.2 | 0.171 | 0.186 |  |  | 3.774 | 3.582 | 3.273 | 3.634 |  |  | 0.16225 | 0.15025 | 0.12125 | 0.13625 | --- | --- | 0.156 | -0.036 | -0.345 | 0.016 | --- | --- | 0.0062255 | -0.0015514 | -0.018423 | 0.00076035 | --- | --- | -0.0032472 | 0.01063 | #NUM! | #NUM! | -5.0791 | #NUM! | #NUM! | -7.1817 | --- | ---
Positive Control (Colony 1) | 0.221 | 0.158 | 0.162 | 0.126 |  |  | 9.619 | 9.845 | 8.415 | 8.481 |  |  | 0.17125 | 0.10825 | 0.11225 | 0.07625 | --- | --- | 6.001 | 6.227 | 4.797 | 4.863 | --- | --- | 0.2269 | 0.37246 | 0.2767 | 0.41295 | --- | --- | 0.32225 | 0.085469 | 0.31348 | 1.3156 | -1.4833 | -0.98762 | -1.2848 | -0.88443 | --- | ---
Positive Control (Colony 2) | 0.155 | 0.131 | 0.129 | 0.094 |  |  | 10.01 | 9.419 | 7.976 | 8.518 |  |  | 0.10525 | 0.08125 | 0.07925 | 0.04425 | --- | --- | 6.392 | 5.801 | 4.358 | 4.9 | --- | --- | 0.39323 | 0.46229 | 0.35606 | 0.71699 | --- | --- | 0.48214 | 0.16264 | 0.46414 | 1.3624 | -0.93336 | -0.77157 | -1.0327 | -0.33269 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.249 | 0.179 | 0.162 | 0.177 |  |  | 7.927 | 7.427 | 7.174 | 6.333 |  |  | 0.19925 | 0.12925 | 0.11225 | 0.12725 | --- | --- | 4.309 | 3.809 | 3.556 | 2.715 | --- | --- | 0.14003 | 0.19081 | 0.20512 | 0.13815 | --- | --- | 0.16853 | 0.034501 | 0.16588 | 1.2283 | -1.9659 | -1.6565 | -1.5842 | -1.9794 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.123 | 0.128 | 0.123 | 0.109 |  |  | 7.626 | 7.287 | 7.072 | 7.061 |  |  | 0.07325 | 0.07825 | 0.07325 | 0.05925 | --- | --- | 4.008 | 3.669 | 3.454 | 3.443 | --- | --- | 0.35428 | 0.3036 | 0.30531 | 0.37625 | --- | --- | 0.33486 | 0.036245 | 0.3334 | 1.1138 | -1.0377 | -1.1921 | -1.1864 | -0.97749 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.176 | 0.175 | 0.159 | 0.183 |  |  | 17.02 | 14.29 | 13.53 | 15.27 |  |  | 0.12625 | 0.12525 | 0.10925 | 0.13325 | --- | --- | 13.402 | 10.672 | 9.912 | 11.652 | --- | --- | 0.68734 | 0.5517 | 0.58745 | 0.56619 | --- | --- | 0.59817 | 0.061232 | 0.59594 | 1.1036 | -0.37493 | -0.59476 | -0.53196 | -0.56882 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.146 | 0.111 | 0.157 | 0.119 |  |  | 12.22 | 11.69 | 11.9 | 11.97 |  |  | 0.09625 | 0.06125 | 0.10725 | 0.06925 | --- | --- | 8.602 | 8.072 | 8.282 | 8.352 | --- | --- | 0.57867 | 0.85331 | 0.5 | 0.78091 | --- | --- | 0.67822 | 0.16621 | 0.66264 | 1.285 | -0.54702 | -0.15863 | -0.69315 | -0.24729 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.382 | 0.249 | 0.257 | 0.229 |  |  | 4.145 | 3.822 | 3.865 | 3.808 |  |  | 0.33225 | 0.19925 | 0.20725 | 0.17925 | --- | --- | 0.527 | 0.204 | 0.247 | 0.19 | --- | --- | 0.01027 | 0.0066292 | 0.0077167 | 0.0068632 | --- | --- | 0.0078698 | 0.0016671 | 0.0077491 | 1.2198 | -4.5785 | -5.0163 | -4.8644 | -4.9816 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.209 | 0.175 | 0.203 | 0.198 |  |  | 3.978 | 3.851 | 3.866 | 3.536 |  |  | 0.15925 | 0.12525 | 0.15325 | 0.14825 | --- | --- | 0.36 | 0.233 | 0.248 | -0.082 | --- | --- | 0.014637 | 0.012045 | 0.010478 | -0.0035814 | --- | --- | 0.0083947 | 0.0081662 | #NUM! | #NUM! | -4.2242 | -4.4191 | -4.5585 | #NUM! | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.204 | 0.149 | 0.181 | 0.127 |  |  | 12.43 | 11.73 | 11.67 | 13.63 |  |  | 0.15425 | 0.09925 | 0.13125 | 0.07725 | --- | --- | 8.812 | 8.112 | 8.052 | 10.012 | --- | --- | 0.3699 | 0.52921 | 0.39722 | 0.83918 | --- | --- | 0.53388 | 0.21509 | 0.50542 | 1.4503 | -0.99453 | -0.63637 | -0.92325 | -0.17533 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.139 | 0.086 | 0.116 | 0.086 |  |  | 10.73 | 10.44 | 10.69 | 12.67 |  |  | 0.08925 | 0.03625 | 0.06625 | 0.03625 | --- | --- | 7.112 | 6.822 | 7.072 | 9.052 | --- | --- | 0.51596 | 1.2185 | 0.69118 | 1.6168 | --- | --- | 1.0106 | 0.50248 | 0.91554 | 1.683 | -0.66173 | 0.19764 | -0.36936 | 0.48048 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.36 | 0.274 | 0.201 | 0.206 |  |  | 6.417 | 5.995 | 6.04 | 6.064 |  |  | 0.31025 | 0.22425 | 0.15125 | 0.15625 | --- | --- | 2.799 | 2.377 | 2.422 | 2.446 | --- | --- | 0.058415 | 0.068632 | 0.10368 | 0.10136 | --- | --- | 0.083023 | 0.022919 | 0.080567 | 1.3311 | -2.8402 | -2.679 | -2.2664 | -2.2891 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.181 | 0.185 | 0.191 | 0.184 |  |  | 6.392 | 6.139 | 6.22 | 6.096 |  |  | 0.13125 | 0.13525 | 0.14125 | 0.13425 | --- | --- | 2.774 | 2.521 | 2.602 | 2.478 | --- | --- | 0.13685 | 0.12069 | 0.11928 | 0.11951 | --- | --- | 0.12408 | 0.0085334 | 0.12387 | 1.0689 | -1.9889 | -2.1145 | -2.1263 | -2.1243 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.206 | 0.225 | 0.181 | 0.181 |  |  | 3.616 | 4.102 | 3.682 | 3.955 |  |  | 0.15625 | 0.17525 | 0.13125 | 0.13125 | --- | --- | -0.002 | 0.484 | 0.064 | 0.337 | --- | --- | -8.2878e-05 | 0.017882 | 0.0031573 | 0.016625 | --- | --- | 0.0093954 | 0.0091841 | #NUM! | #NUM! | #NUM! | -4.024 | -5.758 | -4.0968 | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.158 | 0.132 | 0.162 | 0.125 |  |  | 3.803 | 3.901 | 3.846 | 3.914 |  |  | 0.10825 | 0.08225 | 0.11225 | 0.07525 | --- | --- | 0.185 | 0.283 | 0.228 | 0.296 | --- | --- | 0.011066 | 0.022278 | 0.013152 | 0.025469 | --- | --- | 0.017991 | 0.0069687 | 0.016952 | 1.496 | -4.5039 | -3.8041 | -4.3312 | -3.6703 | --- | ---
Hour 4:
Negative Control (Colony 1) | 0.386 | 0.334 | 0.32 | 0.323 |  |  | 4.106 | 3.966 | 3.965 | 3.796 |  |  | 0.33625 | 0.28425 | 0.27025 | 0.27325 | --- | --- | 0.488 | 0.348 | 0.347 | 0.178 | --- | --- | 0.009397 | 0.007927 | 0.0083137 | 0.0042179 | --- | --- | 0.0074639 | 0.0022517 | 0.007149 | 1.4319 | -4.6674 | -4.8375 | -4.7898 | -5.4684 | --- | ---
Negative Control (Colony 2) | 0.265 | 0.271 | 0.307 | 0.289 |  |  | 3.712 | 3.741 | 3.58 | 3.58 |  |  | 0.21525 | 0.22125 | 0.25725 | 0.23925 | --- | --- | 0.094 | 0.123 | -0.038 | -0.038 | --- | --- | 0.0028276 | 0.0035996 | -0.00095645 | -0.0010284 | --- | --- | 0.0011106 | 0.0024489 | #NUM! | #NUM! | -5.8683 | -5.6269 | #NUM! | #NUM! | --- | ---
Positive Control (Colony 1) | 0.202 | 0.164 | 0.169 | 0.176 |  |  | 15.96 | 14.43 | 15.06 | 16.85 |  |  | 0.15225 | 0.11425 | 0.11925 | 0.12625 | --- | --- | 12.342 | 10.812 | 11.442 | 13.232 | --- | --- | 0.52488 | 0.61275 | 0.62126 | 0.67862 | --- | --- | 0.60938 | 0.063474 | 0.60682 | 1.1127 | -0.64459 | -0.4898 | -0.476 | -0.3877 | --- | ---
Positive Control (Colony 2) | 0.232 | 0.176 | 0.206 | 0.163 |  |  | 15.96 | 16.19 | 16.49 | 16.96 |  |  | 0.18225 | 0.12625 | 0.15625 | 0.11325 | --- | --- | 12.342 | 12.572 | 12.872 | 13.342 | --- | --- | 0.43848 | 0.64477 | 0.53341 | 0.76281 | --- | --- | 0.59487 | 0.14015 | 0.58238 | 1.2701 | -0.82444 | -0.43886 | -0.62847 | -0.27075 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.183 | 0.138 | 0.127 | 0.135 |  |  | 8.389 | 8.387 | 8.374 | 8.47 |  |  | 0.13325 | 0.08825 | 0.07725 | 0.08525 | --- | --- | 4.771 | 4.769 | 4.756 | 4.852 | --- | --- | 0.23183 | 0.3499 | 0.39863 | 0.36852 | --- | --- | 0.33722 | 0.073072 | 0.3304 | 1.274 | -1.4617 | -1.0501 | -0.91971 | -0.99827 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.114 | 0.099 | 0.13 | 0.116 |  |  | 7.627 | 7.939 | 7.63 | 7.988 |  |  | 0.06425 | 0.04925 | 0.08025 | 0.06625 | --- | --- | 4.009 | 4.321 | 4.012 | 4.37 | --- | --- | 0.40401 | 0.56808 | 0.3237 | 0.4271 | --- | --- | 0.43072 | 0.10173 | 0.42206 | 1.2604 | -0.90631 | -0.56549 | -1.1279 | -0.85074 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.238 | 0.193 | 0.221 | 0.22 |  |  | 23.84 | 23.58 | 24.47 | 24.18 |  |  | 0.18825 | 0.14325 | 0.17125 | 0.17025 | --- | --- | 20.222 | 19.962 | 20.852 | 20.562 | --- | --- | 0.69554 | 0.90228 | 0.7884 | 0.78201 | --- | --- | 0.79206 | 0.084812 | 0.78869 | 1.1123 | -0.36307 | -0.10283 | -0.23774 | -0.24589 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.185 | 0.182 | 0.192 | 0.2 |  |  | 19.62 | 18.82 | 19.69 | 20.21 |  |  | 0.13525 | 0.13225 | 0.14225 | 0.15025 | --- | --- | 16.002 | 15.202 | 16.072 | 16.592 | --- | --- | 0.76607 | 0.74428 | 0.73156 | 0.71502 | --- | --- | 0.73923 | 0.021533 | 0.739 | 1.0295 | -0.26648 | -0.29534 | -0.31258 | -0.33545 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.313 | 0.257 | 0.27 | 0.332 |  |  | 4.397 | 4.284 | 4.224 | 4.114 |  |  | 0.26325 | 0.20725 | 0.22025 | 0.28225 | --- | --- | 0.779 | 0.666 | 0.606 | 0.496 | --- | --- | 0.01916 | 0.020807 | 0.017815 | 0.011378 | --- | --- | 0.01729 | 0.0041268 | 0.01686 | 1.3096 | -3.9549 | -3.8725 | -4.0277 | -4.476 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.218 | 0.228 | 0.247 | 0.276 |  |  | 4.426 | 3.922 | 4.078 | 3.93 |  |  | 0.16825 | 0.17825 | 0.19725 | 0.22625 | --- | --- | 0.808 | 0.304 | 0.46 | 0.312 | --- | --- | 0.031095 | 0.011043 | 0.0151 | 0.0089289 | --- | --- | 0.016542 | 0.010034 | 0.014668 | 1.7253 | -3.4707 | -4.506 | -4.1931 | -4.7185 | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.233 | 0.162 | 0.077 | 0.083 |  |  | 12.25 | 11.89 | 1.99 | 11.74 |  |  | 0.18325 | 0.11225 | 0.02725 | 0.03325 | --- | --- | 8.632 | 8.272 | -1.628 | 8.122 | --- | --- | 0.305 | 0.47715 | -0.38683 | 1.5816 | --- | --- | 0.49424 | 0.81543 | #NUM! | #NUM! | -1.1874 | -0.73992 | #NUM! | 0.45845 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.089 | 0.079 | 0.089 | 0.081 |  |  | 11.69 | 11.07 | 12.45 | 12.98 |  |  | 0.03925 | 0.02925 | 0.03925 | 0.03125 | --- | --- | 8.072 | 7.452 | 8.832 | 9.362 | --- | --- | 1.3316 | 1.6496 | 1.457 | 1.9398 | --- | --- | 1.5945 | 0.26475 | 1.5785 | 1.1771 | 0.28638 | 0.50053 | 0.37636 | 0.66257 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.343 | 0.252 | 0.118 | 0.213 |  |  | 7.217 | 7.029 | 6.908 | 6.907 |  |  | 0.29325 | 0.20225 | 0.06825 | 0.16325 | --- | --- | 3.599 | 3.411 | 3.29 | 3.289 | --- | --- | 0.079465 | 0.1092 | 0.31212 | 0.13045 | --- | --- | 0.15781 | 0.10498 | 0.1371 | 1.7959 | -2.5324 | -2.2146 | -1.1644 | -2.0368 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.205 | 0.189 | 0.194 | 0.196 |  |  | 7.026 | 6.85 | 6.502 | 6.833 |  |  | 0.15525 | 0.13925 | 0.14425 | 0.14625 | --- | --- | 3.408 | 3.232 | 2.884 | 3.215 | --- | --- | 0.14213 | 0.15028 | 0.12945 | 0.14234 | --- | --- | 0.14105 | 0.0086133 | 0.14085 | 1.0639 | -1.951 | -1.8952 | -2.0444 | -1.9496 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.266 | 0.253 | 0.232 | 0.236 |  |  | 4.223 | 4.154 | 3.93 | 3.988 |  |  | 0.21625 | 0.20325 | 0.18225 | 0.18625 | --- | --- | 0.605 | 0.536 | 0.312 | 0.37 | --- | --- | 0.018115 | 0.017075 | 0.011085 | 0.012863 | --- | --- | 0.014784 | 0.0033526 | 0.014492 | 1.2623 | -4.011 | -4.0701 | -4.5022 | -4.3534 | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.181 | 0.178 | 0.19 | 0.184 |  |  | 4.245 | 4.192 | 4.024 | 4.266 |  |  | 0.13125 | 0.12825 | 0.14025 | 0.13425 | --- | --- | 0.627 | 0.574 | 0.406 | 0.648 | --- | --- | 0.030931 | 0.028979 | 0.018744 | 0.031253 | --- | --- | 0.027477 | 0.0059082 | 0.026919 | 1.2759 | -3.476 | -3.5412 | -3.9769 | -3.4656 | --- | ---
Hour 6:
Negative Control (Colony 1) | 0.413 | 0.383 | 0.394 | 0.395 |  |  | 4.374 | 4.253 | 4.099 | 4.364 |  |  | 0.36325 | 0.33325 | 0.34425 | 0.34525 | --- | --- | 0.756 | 0.635 | 0.481 | 0.746 | --- | --- | 0.013476 | 0.012338 | 0.009047 | 0.013991 | --- | --- | 0.012213 | 0.0022206 | 0.012044 | 1.2189 | -4.3069 | -4.3951 | -4.7053 | -4.2694 | --- | ---
Negative Control (Colony 2) | 0.362 | 0.341 | 0.334 | 0.349 |  |  | 4.123 | 4.096 | 4.029 | 4.33 |  |  | 0.31225 | 0.29125 | 0.28425 | 0.29925 | --- | --- | 0.505 | 0.478 | 0.411 | 0.712 | --- | --- | 0.010472 | 0.010627 | 0.0093621 | 0.015406 | --- | --- | 0.011467 | 0.0026857 | 0.011256 | 1.242 | -4.5591 | -4.5444 | -4.6711 | -4.173 | --- | ---
Positive Control (Colony 1) | 0.315 | 0.286 | 0.302 | 0.312 |  |  | 29.19 | 24.21 | 28.59 | 27.84 |  |  | 0.26525 | 0.23625 | 0.25225 | 0.26225 | --- | --- | 25.572 | 20.592 | 24.972 | 24.222 | --- | --- | 0.62423 | 0.56436 | 0.64099 | 0.59803 | --- | --- | 0.6069 | 0.033419 | 0.60621 | 1.0572 | -0.47124 | -0.57206 | -0.44474 | -0.51411 | --- | ---
Positive Control (Colony 2) | 0.342 | 0.323 | 0.346 | 0.277 |  |  | 26.78 | 27.11 | 26.46 | 27.21 |  |  | 0.29225 | 0.27325 | 0.29625 | 0.22725 | --- | --- | 23.162 | 23.492 | 22.842 | 23.592 | --- | --- | 0.51316 | 0.55666 | 0.49924 | 0.67219 | --- | --- | 0.56031 | 0.078493 | 0.55643 | 1.1436 | -0.66717 | -0.5858 | -0.69467 | -0.39721 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.341 | 0.238 | 0.21 | 0.208 |  |  | 9.903 | 9.307 | 9.974 | 9.856 |  |  | 0.29125 | 0.18825 | 0.16025 | 0.15825 | --- | --- | 6.285 | 5.689 | 6.356 | 6.238 | --- | --- | 0.13972 | 0.19567 | 0.25681 | 0.25523 | --- | --- | 0.21186 | 0.055879 | 0.20575 | 1.333 | -1.9681 | -1.6313 | -1.3594 | -1.3656 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.176 | 1.159 | 0.156 | 0.143 |  |  | 9.45 | 9.541 | 9.245 | 9.339 |  |  | 0.12625 | 1.1093 | 0.10625 | 0.09325 | --- | --- | 5.832 | 5.923 | 5.627 | 5.721 | --- | --- | 0.2991 | 0.034574 | 0.34291 | 0.39724 | --- | --- | 0.26846 | 0.16101 | 0.19373 | 3.1732 | -1.207 | -3.3647 | -1.0703 | -0.92321 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.294 | 0.361 | 0.289 | 0.331 |  |  | 35.49 | 34.2 | 34.61 | 33.49 |  |  | 0.24425 | 0.31125 | 0.23925 | 0.28125 | --- | --- | 31.872 | 30.582 | 30.992 | 29.872 | --- | --- | 0.8449 | 0.63619 | 0.83874 | 0.68771 | --- | --- | 0.75189 | 0.10599 | 0.7462 | 1.1535 | -0.16853 | -0.45225 | -0.17585 | -0.37439 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.347 | 0.3 | 0.268 | 0.294 |  |  | 27.4 | 27.24 | 26.3 | 29.43 |  |  | 0.29725 | 0.25025 | 0.21825 | 0.24425 | --- | --- | 23.782 | 23.622 | 22.682 | 25.812 | --- | --- | 0.51803 | 0.61119 | 0.67291 | 0.68426 | --- | --- | 0.6216 | 0.076143 | 0.61791 | 1.1361 | -0.65771 | -0.49235 | -0.39614 | -0.37942 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.502 | 0.425 | 0.426 | 0.313 |  |  | 4.552 | 4.491 | 4.471 | 4.705 |  |  | 0.45225 | 0.37525 | 0.37625 | 0.26325 | --- | --- | 0.934 | 0.873 | 0.853 | 1.087 | --- | --- | 0.013372 | 0.015063 | 0.014679 | 0.026736 | --- | --- | 0.017463 | 0.0062243 | 0.016768 | 1.3705 | -4.3146 | -4.1955 | -4.2213 | -3.6218 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.363 | 0.258 | 0.334 | 0.323 |  |  | 4.297 | 4.397 | 4.202 | 4.268 |  |  | 0.31325 | 0.20825 | 0.28425 | 0.27325 | --- | --- | 0.679 | 0.779 | 0.584 | 0.65 | --- | --- | 0.014035 | 0.024221 | 0.013303 | 0.015402 | --- | --- | 0.01674 | 0.0050623 | 0.016245 | 1.314 | -4.2662 | -3.7206 | -4.3198 | -4.1732 | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.24 | 0.191 | 0.171 | 0.151 |  |  | 13.73 | 13.48 | 13.24 | 13.3 |  |  | 0.19025 | 0.14125 | 0.12125 | 0.10125 | --- | --- | 10.112 | 9.862 | 9.622 | 9.682 | --- | --- | 0.34415 | 0.45207 | 0.51383 | 0.61916 | --- | --- | 0.4823 | 0.11507 | 0.47167 | 1.2797 | -1.0667 | -0.79391 | -0.66587 | -0.47939 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.173 | 0.133 | 0.163 | 0.12 |  |  | 16.8 | 16.81 | 16.84 | 16.73 |  |  | 0.12325 | 0.08325 | 0.11325 | 0.07025 | --- | --- | 13.182 | 13.192 | 13.222 | 13.112 | --- | --- | 0.69251 | 1.026 | 0.75595 | 1.2085 | --- | --- | 0.92075 | 0.24024 | 0.8976 | 1.2972 | -0.36743 | 0.025693 | -0.27979 | 0.1894 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.455 | 0.267 | 0.278 | 0.342 |  |  | 7.989 | 8.519 | 8.006 | 7.904 |  |  | 0.40525 | 0.21725 | 0.22825 | 0.29225 | --- | --- | 4.371 | 4.901 | 4.388 | 4.286 | --- | --- | 0.069838 | 0.14607 | 0.12448 | 0.094958 | --- | --- | 0.10884 | 0.033388 | 0.10479 | 1.3822 | -2.6616 | -1.9237 | -2.0836 | -2.3543 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.328 | 0.288 | 0.276 | 0.299 |  |  | 7.333 | 7.335 | 7.997 | 7.476 |  |  | 0.27825 | 0.23825 | 0.22625 | 0.24925 | --- | --- | 3.715 | 3.717 | 4.379 | 3.858 | --- | --- | 0.086448 | 0.10102 | 0.12532 | 0.10022 | --- | --- | 0.10325 | 0.016161 | 0.10234 | 1.1652 | -2.4482 | -2.2925 | -2.0769 | -2.3004 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.365 | 0.364 | 0.36 | 0.339 |  |  | 4.282 | 4.116 | 4.094 | 4.008 |  |  | 0.31525 | 0.31425 | 0.31025 | 0.28925 | --- | --- | 0.664 | 0.498 | 0.476 | 0.39 | --- | --- | 0.013638 | 0.010261 | 0.0099341 | 0.0087302 | --- | --- | 0.010641 | 0.0021037 | 0.010496 | 1.2068 | -4.2949 | -4.5794 | -4.6118 | -4.741 | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.317 | 0.298 | 0.318 | 0.289 |  |  | 4.113 | 4.12 | 4.043 | 4.578 |  |  | 0.26725 | 0.24825 | 0.26825 | 0.23925 | --- | --- | 0.495 | 0.502 | 0.425 | 0.96 | --- | --- | 0.011993 | 0.013093 | 0.010258 | 0.025981 | --- | --- | 0.015331 | 0.0071949 | 0.014303 | 1.5076 | -4.4235 | -4.3357 | -4.5797 | -3.6504 | --- | ---